**Narragansett Bay (NAR) NERR Nutrient Metadata**

**January 2012 – December 2012**

**Latest Update: October 15, 2015**

# I. Data Set and Research Descriptors

1. **Principal investigator(s) and contact persons**
   1. **Reserve Contacts**

55 South Reserve Drive

Prudence Island, RI 02872

Fax: 401-682-7366

Kenneth Raposa, PhD

Research Coordinator

(401) 683-7849

[kenny@nbnerr.org](mailto:Kenny@nbnerr.org)

Daisy Durant, PhD

Marine Research Specialist II &

SWMP Technician

Responsible for data management

(401) 683-7368

[daisy@nbnerr.org](mailto:chris@nbnerr.org)

* 1. **Laboratory Contacts**

Nutrient analyses

Carol Pollard, Lab Manager

Virginia Institute of Marine Science

Analytical Service Center for Nutrients

PO Box 1346

Goucester Point, VA 23062

Phone: (804) 684-7213

Email: [pollard@vims.edu](mailto:pollard@vims.edu)

Chlorophyll analyses

Jeffrey Mercer/Laura Reed

Lab Technicians

Marine Ecosystem Research Laboratory (MERL)

University of Rhode Island

Graduate School of Oceanography

Narragansett, Rhode Island 02882

Phone: (401) 874-6619

E-mail: [jmercer@gso.uri.edu](mailto:jmercer@gso.uri.edu)

[reedlhht@aol.com](mailto:reedlhht@aol.com)

**2) Research Objectives**

Nutrient and chlorophyll samples are being collected off Prudence Island in Narragansett Bay as part of the National Estuarine Research Reserve's (NERR) System-Wide Monitoring Program (SWMP). The goal is to develop long-term datasets for representative estuarine systems in order to track changes in water quality over time. Because Prudence Island is located in the geographic center of Narragansett Bay, it is an ideal location for monitoring the status and trends in water quality in the Bay over time. One NERR monitoring station has been established at Potter Cove, on the island's northeastern shore. This area is impacted by boat traffic and storm runoff from mainland urban and residential areas. Two NERR monitoring stations are located at T-Wharf, on the southeastern shore of the island facing the open waters of Rhode Island Sound. The two T-Wharf stations (T-Wharf Surface and T-Wharf Bottom) measure conditions in the upper (~0.5 m from the surface) and bottom (~1 m from the bottom) layers of the water column, respectively. These stations are approximately 10 km south of the Potter Cove site. Boat traffic is sparse at T-Wharf and storm runoff is less likely to have a significant impact on water quality. A fourth monitoring site is located in Nag Creek, a salt marsh tidal creek which flows into the West Passage of Narragansett Bay. The addition of this site completes our representation of dominant habitat types occurring in Narragansett Bay (i.e. marsh, cove, and open water).

* 1. **Monthly Grab Sampling Program**

Monthly grab samples are collected once per month from each of the four stations to quantify seasonal patterns of selected nutrient species and chlorophyll in different estuarine habitats (marsh creek, cove, open water surface, open water bottom).

* 1. **Diel Sampling Program**

Once per month, samples are collected at approximately 2 hour 15 minute intervals over a 24 hour cycle (lunar day) to document changes in nutrients and chlorophyll in response to tidal forcing. Previously (from 2002 to 2010) diel sampling occurred at T-Wharf. However, after analyzing the historic data from the site, no significant trend or patterns were found over time. Thus, the station was moved and has been at Potter Cove since January of 2011, in order to characterize nutrients and chlorophyll from this site.

1. **Research methods**
   1. **Monthly Grab Sampling Program**

Monthly grab samples were taken at the four permanent stations (Potter Cove, Nag Creek, T-Wharf Surface, and T-Wharf Bottom) each month throughout 2012.

All grab samples were taken on the same day between +3 hrs before slack low-water and slack low-water. No distinction was made between neap and spring tide conditions. Replicate (N=2) samples were collected by hand at Nag Creek and with a horizontal Van-Dorn sampling bottle at the other three sites. At Nag Creek, samples were collected at approximately 0.25 m below the surface, at Potter Cove and T-Wharf Bottom at 1 m off the bottom, and at T-Wharf Surface at 1 m below the surface. In the field, all samples were transferred to Nalgene amber wide-mouth HDPE 500 ml sample bottles. All sampling bottles were previously acid washed in a solution of 10 % Hydrochloric acid, rinsed (3x) with distilled-deionized water, and let dry. In the field, the bottles were rinsed (3x) with ambient water prior to collection of the sample. Samples were immediately placed on ice and in the dark and returned to the Reserve’s laboratory for filtering the same day of collection.

For pigment analysis (chlorophyll and phaeophytin), we filtered 20 ml of sample onto a glass microfiber filter Type GF/F (Whatman 1825-025), added 2 drops of magnesium carbonate (MgCO3) solution to the filter, folded in half, wrapped in aluminum foil, and stored at -20oC until transported to the analytical lab.


For nutrient analyses (ammonia, nitrate-nitrite, nitrite, phosphate, and silicate) we filtered 60 ml using a nucleopore polycarbonate Track-Etch membrane 0.4 µm pore size (Whatman 111107) into a small sampling bottle which was previously acid washed, rinsed and dried as described above. After filtering, all samples were frozen at -20oC until transported to the analytical lab.

Filtered samples were transported to the Marine Ecosystems Research Laboratory (MERL) at the University of Rhode Island Bay Campus, Graduate School of Oceanography (URI/GSO) for chlorophyll and phaeophytin pigment analyses, and shipped overnight to the Virginia Institute of Marine Science lab for nutrient analyses.

* 1. **Diel Sampling Program**

The diel sampling program was conducted once per month throughout 2012. For this program, an ISCO 6712 portable water sampler was deployed at the Potter Cove station. For this program, an ISCO 6712 portable water sampler was deployed at the T-Wharf Bottom station. This device was programed to automatically sample 500 ml of water at 2 hr 15 min intervals, covering a complete semidiurnal tidal cycle. All samples were pumped into polyethylene sample bottles that were previously acid washed in a solution of 10 % Hydrochloric acid, rinsed (3x) with distilled-deionized water, and let dry. A total of 12 samples were collected each month. During the warm months (usually May to October), ice gel packs were used inside the sampler to keep the samples cool during the sampling period. At the end of the 24 hr. period, the 12 samples were kept in the dark and returned to the laboratory for filtering the same day. Samples were filtered and frozen as described in Section (a) above. The samples were transported to the same laboratories mentioned above for chlorophyll and phaeophytin pigment analyses and nutrient analyses.

1. **Site Location and Character**

The NBNERR consists of 4453 acres of diverse estuarine and terrestrial habitats ranging from deep water to salt marshes to forested uplands. The land holdings include approximately 65% of Prudence Island, most of nearby Patience Island, and all of Hope and Dyer Islands. The Reserve is located close to the geographic center of Narragansett Bay in Rhode Island. The Bay has a drainage basin of 1800 square miles and there are numerous and substantial freshwater inputs to the Bay

The Narragansett Bay watershed consists of nine subwatersheds draining an area of approximately 4,836square km [[1]](#footnote-1)(Pilson, 1985) and numerous and substantial freshwater inputs to the Bay. Approximately 39% of the watershed lies in Rhode Island and 61% in Massachusetts. It is referred to as a shallow estuary; however, its water depth varies considerably. Depth averages approximately 9.0 m throughout the Bay, but it’s deeper in the East Passage (approximately 15.2 m) and shallower in the West Passage (approximately 7.5 m). More information and a detailed description of the Narragansett Bay NERR and the Narragansett Bay watershed can be found in 2Raposa and Schwartz, available at <http://www.nbnerr.org/profile.htm>.

Specific characteristics of the Narragansett Bay National Estuarine Research Reserve

Location: 41°38’30” N, 71°20’30” W

Tidal range: -0.2 to 1.7 meters MLW

Salinity: 15 to 32 pp

Temperature: -1.0 to 26 C

Province: North temperate, Virginian bioregion

Specific characteristics of the Potter Cove site are:

Location: 41o38’25.98 N, 71o20’ 27.17 W

Depth: 0.9 to 3.9 meters

Bottom habitat: Sand, silt, some organic mud

Pollutants: Boaters’ wastes, storm runoff from mainland urban areas

Watershed: Narragansett Bay, North Prudence (4801 square km)

Specific characteristics of the Nag Creek site are:

Location: 41o37’29.46 N, 71o19 27.42 W

Depth: 0.1 to 1.4 meters

Bottom habitat: Organic mud

Pollutants: Negligible

Watershed: Narragansett Bay, West Passage

Specific characteristics of the T-Wharf Surface site are:

Location: 41o34’42.10 N, 71o19’16.05 W

Depth: 0.2 to 0.9 meters

Bottom habitat: Sand, silt, some organic mud

Pollutants: Negligible

Watershed: Narragansett Bay, South Prudence

Specific characteristics of the T-Wharf Bottom site are:

Location: 41o34’42.10 N, 71o19’16.05 W

Depth: 4.6 to 6.9 meters

Bottom habitat: Sand, silt, some organic mud

Pollutants: Negligible

Watershed: Narragansett Bay, South Prudence

1. **Coded Variable Definitions**

Station Code Names

narncnut = Nag Creek

narpcnut = Potter Cove

nartbnut = T-Wharf Bottom

nartsnut = T-Wharf Surface

Monthly grab sample program = 1

Diel grab sample program = 2

1. **Data collection period**

The Narragansett Bay National Estuarine Research Reserve started grab sampling started at each of the four sites in the spring of 2002. Diel sampling started in 2002 at T-Wharf Bottom until 2010; it has been at Potter Cove since 2011. For both Grab and Diel sampling, time is coded based on a 2400 hour clock and is referenced to Standard Time (ST) format.

Monthly Grab Sampling

|  |  |  |  |
| --- | --- | --- | --- |
| Site | Start Date | Start Time | End Time |
| narncnut | 01/09/12 | 12:20 | 12:23 |
| narncnut | 02/09/12 | 13:20 | 13:23 |
| narncnut | 03/07/12 | 11:24 | 11:25 |
| narncnut | 04/04/12 | 09:34 | 09:35 |
| narncnut | 05/02/12 | 08:35 | 08:36 |
| narncnut | 06/06/12 | 14:30 | 14:31 |
| narncnut | 07/08/12 | 15:10 | 15:13 |
| narncnut | 08/16/12 | 11:18 | 11:19 |
| narncnut | 09/14/12 | 10:50 | 10:51 |
| narncnut | 10/12/12 | 09:50 | 09:53 |
| narncnut | 11/13/12 | 11:26 | 11:28 |
| narncnut | 12/03/12 | 14:20 | 14:22 |
| narpcnut | 01/09/12 | 12:02 | 12:05 |
| narpcnut | 02/09/12 | 12:20 | 12:23 |
| narpcnut | 03/07/12 | 11:00 | 11:03 |
| narpcnut | 04/04/12 | 09:10 | 09:13 |
| narpcnut | 05/02/12 | 08:15 | 08:18 |
| narpcnut | 06/06/12 | 14:05 | 14:08 |
| narpcnut | 07/08/12 | 14:50 | 14:53 |
| narpcnut | 08/16/12 | 11:45 | 11:48 |
| narpcnut | 09/14/12 | 10:20 | 10:23 |
| narpcnut | 10/12/12 | 09:20 | 09:22 |
| narpcnut | 11/13/12 | 11:42 | 11:44 |
| narpcnut | 12/03/12 | 14:00 | 14:02 |

Monthly Grab Sampling (cont.)

|  |  |  |  |
| --- | --- | --- | --- |
| Site | Start Date | Start Time | End Time |
| nartbnut | 01/09/12 | 13:02 | 13:05 |
| nartbnut | 02/09/12 | 13:55 | 13:58 |
| nartbnut | 03/07/12 | 12:05 | 12:08 |
| nartbnut | 04/04/12 | 10:15 | 10:18 |
| nartbnut | 05/02/12 | 09:05 | 09:08 |
| nartbnut | 06/06/12 | 14:30 | 14:33 |
| nartbnut | 07/08/12 | 15:45 | 15:48 |
| nartbnut | 08/16/12 | 10:50 | 10:53 |
| nartbnut | 09/14/12 | 11:20 | 11:23 |
| nartbnut | 10/12/12 | 10:10 | 10:12 |
| nartbnut | 11/13/12 | 11:00 | 11:02 |
| nartbnut | 12/03/12 | 14:45 | 14:47 |
| nartsnut | 01/09/12 | 12:55 | 12:58 |
| nartsnut | 02/09/12 | 14:00 | 14:03 |
| nartsnut | 03/07/12 | 12:00 | 12:04 |
| nartsnut | 04/04/12 | 10:20 | 10:22 |
| nartsnut | 05/02/12 | 09:10 | 09:13 |
| nartsnut | 06/06/12 | 14:35 | 14:38 |
| nartsnut | 07/08/12 | 15:50 | 15:52 |
| nartsnut | 08/16/12 | 10:55 | 10:57 |
| nartsnut | 09/14/12 | 11:28 | 11:30 |
| nartsnut | 10/12/12 | 10:15 | 10:17 |
| nartsnut | 11/13/12 | 10:55 | 10:58 |
| nartsnut | 12/03/12 | 14:40 | 14:42 |

Diel Sampling

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Site | Start Date | Start Time | End Date | End Time |
| narpcnut | 01/25/12 | 12:20 | 01/26/12 | 13:05 |
| narpcnut | 02/23/12 | 11:45 | 02/24/12 | 12:30 |
| narpcnut | 03/21/12 | 10:05 | 03/22/12 | 10:50 |
| narpcnut | 04/17/12 | 08:31 | 04/18/12 | 09:16 |
| narpcnut | 05/23/12 | 12:00 | 05/24/12 | 12:45 |
| narpcnut | 06/19/12 | 10:24 | 06/20/12 | 11:09 |
| narpcnut | 07/30/12 | 08:04 | 07/31/12 | 08:49 |
| narpcnut | 08/27/12 | 07:07 | 08/28/12 | 07:52 |
| narpcnut | 09/26/12 | 08:26 | 09/27/12 | 09:11 |
| narpcnut | 10/24/12 | 07:35 | 10/25/12 | 08:20 |
| narpcnut | 11/27/12 | 10:15 | 11/28/12 | 11:00 |
| narpcnut | 12/16/12 | 00:35 | 12/17/12 | 01:20 |

1. **Associated researchers and projects**

This SWMP nutrient monitoring program is part of a larger, multi-institution effort to monitor water quality and other associated parameters in Narragansett Bay, RI called the Narragansett Bay Window. This includes monitoring efforts conducted by the University of Rhode Island’s Graduate School of Oceanography, the Narragansett Bay Estuary Program, and NMFS and EPA in Narragansett, RI, among others.

In addition, we are implementing a water quality monitoring program through SWMP since 1995. YSI 6600 data loggers are use to collect physical and chemical parameters at 15-minutes intervals at the same four stations we collect samples for the nutrient monitoring program. Water quality data is available near-real time from the T-Wharf Bottom monitoring station through the CDMO website <http://cdmo.baruch.sc.edu/>; historical data from all the monitoring stations are also available through CDMO, NBNERR <http://nbnerr.org/>, or directly contacting the Research Coordinator.

To compliment the water quality and nutrient monitoring programs, and the biological monitoring and research at NBNERR and at Narragansett Bay, NBNERR operates and maintains a weather station that continually collects air temperature, relative humidity, barometric pressure, wind speed and direction, photosynthetic active radiation, and precipitation. All this information is available near-real time through the CDMO website <http://cdmo.baruch.sc.edu/>; historical data is also available through CDMO, NBNERR <http://nbnerr.org/>, or directly contacting the Research Coordinator.

1. **Distribution**

NOAA retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data. The NERRS retains the right to be fully credited for having collected and process the data. Following academic courtesy standards, the NERR site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used. The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement. The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons. The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

Requested citation format:

National Estuarine Research Reserve System (NERRS). 2012. System-wide Monitoring Program. Data accessed from the NOAA NERRS Centralized Data Management Office website: www.nerrsdata.org; accessed 12 October 2012.

NERR nutrient data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal investigators and contact persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page www.nerrsdata.org. Data are available in comma separated version format.

**II. Physical Structure Descriptors**

1. **Entry verification**

Nutrient data are entered into a Microsoft Excel worksheet and processed using the NutrientQAQC Excel macro. The NutrientQAQC macro sets up the data worksheet, metadata worksheets, and MDL worksheet; adds chosen parameters and facilitates data entry; allows the user to set the number of significant figures to be reported for each parameter and rounds using banker’s rounding rules; allows the user to input MDL values and then automatically flags/codes measured values below MDL and inserts the MDL; calculates parameters chosen by the user and automatically flags/codes for component values below MDL, negative calculated values, and missing data; allows the user to apply QAQC flags and codes to the data; produces summary statistics; graphs selected parameters for review; and exports the resulting data file to the CDMO for tertiary QAQC and assimilation into the CDMO’s authoritative online database

The Virginia Institute of Marine Science (VIMS) entered the data into Microsoft Excel spreadsheets, calculated and reported results in mg L-1, and sent the files electronically to the Reserve. For purposes of consistency in the NERR System, concentrations are calculated as mg L-1 based on atomic weights of 14.010, 30.97, and 28.09, and 12.01 for N, P, Si, and C, respectively.

Data entry verification was completed by Dr. Daisy Durant. Final verification and this metadata documentation were checked by Dr. Kenneth Raposa, before being sent to the CDMO permanent database.

**10) Parameter Titles and Variable Names by Category**

Required NOAA/NERRS System-wide Monitoring Program nutrient parameters are denoted by an asterisks “\*”.

|  |  |  |  |
| --- | --- | --- | --- |
| Data Category | Parameter | Variable Name | Units of Measure |
| Phosphorus and | \*Orthophosphate | PO4F | mg/L as P |
| Nitrogen: | \*Ammonium, Filtered | NH4F | mg/L as N |
|  | \*Nitrite, Filtered | NO2F | mg/L as N |
|  | \*Nitrate, Filtered | NO3F | mg/L as N |
|  | \*Nitrite + Nitrate, Filtered | NO23F | mg/L as N |
|  | Dissolved Inorganic Nitrogen | DIN | mg/L as N |
| Plant Pigments: | \*Chlorophyll a | CHLA\_N | µg/L |
|  | Phaeophytin | PHEA | µg/L |
| Other Lab Parameters: | Silicate, Filtered | SiO4F | mg/L as Si |

Notes:

* Time is coded based on a 2400 clock and is referenced to Standard Time format.
* Reserves have the option of measuring either NO2 or NO3 or they may substitute NO23 for individual analyses if they can show that NO2 is a minor component relative to NO3.

**11) Measured and Calculated Laboratory Parameters**

1. **Parameters Measured Directly**

* Nitrogen species: NH4F, NO2F, NO23F
* Phosphorus species: PO4F
* Other: CHLA\_N, PHEA, SiO4F, WTEM

1. **Calculated Parameters**

* NO3F: NO23F-NO2F
* DIN: NO23F+NH4F

**12)** **Limits of Detection**

Method Detection Limits (MDL), the lowest concentration of a parameter that an analytical procedure can reliably detect, have been established by VIMS Analytical Service Center for Nutrient. The MDL is determined as 3 times the standard deviation of a minimum of 7 replicates of a single low concentration sample. These values are reviewed and revised periodically.

Method Detection Limits

|  |  |  |  |
| --- | --- | --- | --- |
| Parameter | Start Date | End Date | MDL |
| PO4F | 01/01/2012 | 12/31/2012 | 0.0020 mg L-1 |
| NH4F | 01/01/2012 | 12/31/2012 | 0.0056 mg L-1 |
| NO2F | 01/01/2012 | 12/31/2012 | 0.0016 mg L-1 |
| NO23F | 01/01/2012 | 12/31/2012 | 0.0047 mg L-1 |
| SiO4F | 01/01/2012 | 12/31/2012 | 0.0800 mg L-1 |
| CHLA\_N | 01/01/2012 | 12/31/2012 | 0.05 µg L-1 |
| PHEA | 01/01/2012 | 12/31/2012 | 0.05 µg L-1 |

**13)** **Laboratory Methods**

1. **Parameter:** **Chlorophyll *a* (CHLA\_N), Phaeophytin (PHEA)**
   * Marine Ecosystems Research Laboratory Method
   * Method References: Oviatt, C. A., and K. M. Hindle, 1994. Manual of biological and geochemical techniques in coastal areas. pp. 3-7.
   * Method Descriptor: Chlorophyll *a* is extracted in 10 ml 90% acetone and fluorescence is measured and recorded (Fo). Two drops of 10% hydrochloric acid are added to convert the chl to phaeopigments (Phae). The fluorescence is again measured and recorded (Fa). The concentration (μg L-1) of Chl *a* and Phae are calculated using the Fo/Fa ratio.
   * Preservation Method: 20 ml of sample is filtered onto a Whatman® Glass Microfiber Filter: Type GF/F (25mm), treated with 2 drops of MgCO3 solution, folded in half, wrapped in aluminum foil, and stored at –4 oC until analysis.
2. **Parameter: NH4F**
   * VIMS Laboratory Method: SA156-350.1 - NH3
   * Method Reference:

* U.S. EPA. 1974. Methods for Chemical Analysis of Water and Wastes, pp. 168-174.
* Standard Methods for the Examination of Water and Wastewater, 14th edition. p 410. Method 418A and 418B (1975).
* Annual Book of ASTM Standards, Part 31. "Water", Standard 1426-74, Method A, p 237 (1976).
* EPA 600/R-97/072 Method 349.0. Determination of Ammonia in Estuarine and Coastal Waters by Gas Segmented Continuous Flow Colorimetric Analysis. IN: Methods for the Determination of Chemical Substances in Marine and Estuarine Environmental Matrices -2nd Edition. National Exposure Research Laboratory, Office of Research and Development U.S. EPA, Cincinnati, Ohio 45268.
  + Method Descriptor: Ammonia reacts with alkaline phenol and hypochlorite to form indophenol blue in an amount that is proportional to the ammonia concentration. The blue color is intensified with sodium nitroprusside. The reaction is catalyzed by heat at 37oC. The range is 0.01 – 2.0 mg L-1.
  + Preservation Method: 40 ml of sample is filtered through a Whatman® Nuclepore Track-Etched Membrane, 0.4 µm disposable disk filter (47mm), and stored at –4 oC until analyzed.

1. **Parameter: NO2F, and NO2F + NO3F**
   * VIMS Laboratory Method: SA461-353.2 - NO2+3, NO2, TDN
   * Method Reference:

* U.S. EPA. 1974. Methods for Chemical Analysis of Water and Wastes, pp. 207 -212.
* Wood, E.D., F.A.G. Armstrong, and F.A. Richards. 1967. Determination of nitrate in seawater by cadmium-copper reduction to nitrite. J. Mar. Biol. Assoc. U.K. 47: 23
* Grasshoff, K., M. Ehrhardt and K. Kremling. 1983. Methods of Seawater Analysis. Verlag Chemie, Federal Republic of Germany. 419 pp.
* EPA 600/R-97/072 Method 353.4. Determination of Nitrate and Nitrite in Estuarine and Coastal Waters by Gas Segmented Flow Colorimetric Analysis. IN: Methods for the Determination of Chemical Substances in Marine and Estuarine Environmental Matrices -2nd Edition. National Exposure Research Laboratory, Office of Research and Development U.S. EPA, Cincinnati, Ohio 45268.
* Method Descriptor: Nitrate is reduced quantitatively to nitrite by cadmium metal in the form of an open tubular reactor. The nitrite thus formed plus any originally present in the sample is colorimetrically detected at 540 nm, following its diazotization with sulfanilamide, and subsequent coupling with N-1-naphthylethylenediamine dihydrochloride. The dissolved nitrate concentrations of the samples are calculated as NO3-N.
  + Preservation Method: 40 ml of sample is filtered through a Whatman® Nuclepore Track-Etched Membrane, 0.4 µm disposable disk filter (47mm), and stored at –4 oC until analyzed.

1. **Parameter:** **SiO4F**
   * VIMS Laboratory Method: Skalar Method Silicates, Catnr. 563-052 issue 101899/MH/99208255
   * Method Reference:

* U.S. EPA. 1974. Methods for Chemical Analysis of Water and Wastes, Method 370.1
* Grasshoff, K., M. Ehrhardt and K. Kremling. 1983. Methods of Seawater Analysis. Verlag Chemie, Federal Republic of Germany. Pp 374-376.
* EPA 600/R-97/072 Method 366.0 Determination of Dissolved Silicate in Estuarine and Coastal Waters by Gas Segmented Flow Colorimetric Analysis. IN: Methods for the Determination of Chemical Substances in Marine and Estuarine Environmental Matrices -2nd Edition. National Exposure Research Laboratory, Office of Research and Development U.S. EPA, Cincinnati, Ohio 45268.
  + Method Descriptor: Based on the reduction of silicomolybdate in acidic solution to “molybedenum blue” by ascorbic acid. Oxalic acid is introduced to the sample stream before the addition of ascorbic acid to eliminate interference from phosphates. The range is 0-1.1 mg Si L-1.
  + Preservation Method: Preservation Method: 40 ml of sample is filtered through a Whatman® Nuclepore Track-Etched Membrane, 0.4 µm disposable disk filter (47mm), and stored at –4 oC until analyzed.

1. **Parameter: PO4F**

* VIMS Laboratory Method SKALAR Method: O-Phosphate / Total Phosphate, Catnr. 503-365.1, issue 042993/MH/93-Demo1
* Method Reference:
  + - Murphy, J. and J.P. Riley. 1962. A modified single solution method for the determination of phosphate in natural waters. Analytica Chim. Acta 27 : 31-36
    - EPA 600/R-97/072 Method 365.5 Determination of Orthophosphate in Estuarine and Coastal Waters by Automated Colorimetric Analysis. IN: Methods for the Determination of Chemical Substances in Marine and Estuarine Environmental Matrices -2nd Edition. National Exposure Research Laboratory, Office of Research and Development. U.S. EPA, Cincinnati, Ohio 45268
  + Method Descriptor: Ammonium molybdate and antimony potassium tartrate react in a sulfuric acid environment to form an antimony-phospho-molybdo complex, which is reduced to a blue colored complex by ascorbic acid. Reaction is heat catalyzed at 40oC.
  + Range is 1 - 50 ppb.
  + Preservation Method: Preservation Method: 40 ml of sample is filtered through a Whatman® Nuclepore Track-Etched Membrane, 0.4 µm disposable disk filter (47mm), and stored at –4 oC until analyzed.

**14) Field and Laboratory QA/QC Programs**

1. **Precision**
   1. **Field Variability** - For the monthly grab sampling program, NBNERR collects two successive grab samples for the determination of water mass variability.
   2. **Laboratory Variability** – replicates of all nutrient samples are run
   3. **Inter-organizational splits** – none
2. **Accuracy**
   1. **Sample Spikes** – blanks
   2. **Standard Reference Material Analysis** – none
   3. **Cross Calibration Exercises** – A cross calibration of the Turner Designs Fluorometer was performed with Battelle Ocean Sciences.

**15) QAQC flag definitions**

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F\_). QAQC flags are applied to the nutrient data during secondary QAQC to indicate data that are out of sensor range low (-4), rejected due to QAQC checks (-3), missing (-2), optional and were not collected (-1), suspect (1), and that have been corrected (5). All remaining data are flagged as having passed initial QAQC checks (0) when the data are uploaded and assimilated into the CDMO ODIS as provisional plus data. The historical data flag (4) is used to indicate data that were submitted to the CDMO prior to the initiation of secondary QAQC flags and codes (and the use of the automated primary QAQC system for WQ and MET data). This flag is only present in historical data that are exported from the CDMO ODIS.

-4 Outside Low Sensor Range

-3 Data Rejected due to QAQC

-2 Missing Data

-1 Optional SWMP Supported Parameter

0 Data Passed Initial QAQC Checks

1 Suspect Data

4 Historical Data: Pre-Auto QAQC

5 Corrected Data

**16) QAQC code definitions**

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column. There are three (3) different code categories, general, sensor, and comment. General errors document general problems with the sample or sample collection, sensor errors document common sensor or parameter specific problems, and comment codes are used to further document conditions or a problem with the data. Only one general or sensor error and one comment code can be applied to a particular data point. However, a record flag column (F\_Record) in the nutrient data allows multiple comment codes to be applied to the entire data record.

General errors

GCM Calculated value could not be determined due to missing data

GCR Calculated value could not be determined due to rejected data

GDM Data missing or sample never collected

GQD Data rejected due to QA/QC checks

GQS Data suspect due to QA/QC checks

Sensor errors

SBL Value below minimum limit of method detection

SCB Calculated value could not be determined due to a below MDL component

SCC Calculation with this component resulted in a negative value

SNV Calculated value is negative

SRD Replicate values differ substantially

SUL Value above upper limit of method detection

Parameter Comments

CAB Algal bloom

CDR Sample diluted and rerun

CHB Sample held beyond specified holding time

CIP Ice present in sample vicinity

CIF Flotsam present in sample vicinity

CLE Sample collected later/earlier than scheduled

CRE Significant rain event

CSM See metadata

CUS Lab analysis from unpreserved sample

Record comments

CAB Algal bloom

CHB Sample held beyond specified holding time

CIP Ice present in sample vicinity

CIF Flotsam present in sample vicinity

CLE Sample collected later/earlier than scheduled

CRE Significant rain event

CSM See metadata

CUS Lab analysis from unpreserved sample

*Cloud cover*

CCL clear (0-10%)

CSP scattered to partly cloudy (10-50%)

CPB partly to broken (50-90%)

COC overcast (>90%)

CFY foggy

CHY hazy

CCC cloud (no percentage)

*Precipitation*

PNP none

PDR drizzle

PLR light rain

PHR heavy rain

PSQ squally

PFQ frozen precipitation (sleet/snow/freezing rain)

PSR mixed rain and snow

*Tide stage*

TSE ebb tide

TSF flood tide

TSH high tide

TSL low tide

*Wave height*

WH0 0 to <0.1 meters

WH1 0.1 to 0.3 meters

WH2 0.3 to 0.6 meters

WH3 0.6 to > 1.0 meters

WH4 1.0 to 1.3 meters

WH5 1.3 or greater meters

*Wind direction*

N from the north

NNE from the north northeast

NE from the northeast

ENE from the east northeast

E from the east

ESE from the east southeast

SE from the southeast

SSE from the south southeast

S from the south

SSW from the south southwest

SW from the southwest

WSW from the west southwest

W from the west

WNW from the west northwest

NW from the northwest

NNW from the north northwest

*Wind speed*

WS0 0 to 1 knot

WS1 > 1 to 10 knots

WS2 > 10 to 20 knots

WS3 > 20 to 30 knots

WS4 > 30 to 40 knots

WS5 > 40 knots

**17)** **Other Remarks**

Data may be missing due to problems with sample collection or processing. Laboratories in the NERRS System submit data that are censored at a lower detection rate limit, called the Method Detection Limit or MDL. MDLs for specific parameters are listed in the Laboratory Methods and Detection Limits Section (Section II, Part 12) of this document. Concentrations that are less than this limit are censored with the use of a QAQC flag and code, and the reported value is the method detection limit itself rather than a measured value. For example, if the measured concentration of NO23F was 0.0005 mg/l as N (MDL=0.0008), the reported value would be 0.0008 and would be flagged as out of sensor range low (-4) and coded SBL. In addition, if any of the components used to calculate a variable are below the MDL, the calculated variable is removed and flagged/coded -4 SCB. If a calculated value is negative, it is rejected and all measured components are marked suspect. If additional information on MDL’s or missing, suspect, or rejected data is needed, contact the Research Coordinator at the Reserve submitting the data.

Note: The way below MDL values are handled in the NERRS SWMP dataset was changed in November of 2011.  Previously, below MDL data from 2007-2010 were also flagged/coded, but either reported as the measured value or a blank cell.  Any 2007-2011 nutrient/pigment data downloaded from the CDMO prior to November of 2011 will reflect this difference.

**Suspect data due to hold times:**

Sample hold times (days) calculated for all grab and diel sampling at NBNERR for samples collected during 2012. **All** but the 7/8 CHLA/PHEA samples are known to have been held significantly longer than allowed per NERRS protocols for **all parameters**. Extended hold times were due to numerous difficulties securing a purchase order through the state of RI. Results should be considered suspect and used with caution.

|  |  |  |  |
| --- | --- | --- | --- |
| Grab Sampling Date | NH4, NO23, NO2, PO4 | SiO4 | CHLA, PHEA |
| 01/09/12 | 386 | 392 | 116 |
| 02/09/12 | 355 | 361 | 85 |
| 03/07/12 | 329 | 335 | 69 |
| 04/04/12 | 303 | 308 | 112 - 114 |
| 05/02/12 | 275 | 280 | 84 - 86 |
| 06/06/12 | 292 | 293 | 49- 51 |
| 07/08/12 | 269 | 274 | 17 - 19 |
| 08/16/12 | 230 | 235 | 168 |
| 09/14/12 | 202 | 207 | 145 |
| 10/12/12 | 167 | 179 | 117 |
| 11/13/12 | 135 | 147 | 85 |
| 12/03/12 | 135 | 136 | 73 |

|  |  |  |  |
| --- | --- | --- | --- |
| Diel Sampling Date | NH4, NO23, NO2, PO4 | SiO4 | CHLA, PHEA |
| 01/25/12 | 370 | 376 | 99 |
| 02/23/12 | 342 | 348 | 82 |
| 03/21/12 | 315 | 321 | 55 |
| 04/17/12 | 290 | 295 | 99 - 101 |
| 05/23/12 | 306 | 307 | 63 - 65 |
| 06/19/12 | 279 | 280 | 36 - 38 |
| 07/30/12 | 247 | 252 | 185 |
| 08/27/12 | 220 | 225 | 157 |
| 09/26/12 | 190 | 195 | 133 |
| 10/24/12 | 174 | 181 | 105 |

Rainfall totals from the NBNERR weather station by month for January – December 2011 are as follows:

|  |  |  |
| --- | --- | --- |
| Month | Monthly Total (mm) | Monthly total (inches) |
| January | 48.3 | 1.9 |
| February | 122.4 | 4.8 |
| March | 59.7 | 2.4 |
| April | 135.6 | 5.3 |
| May | 76.2 | 3.0 |
| June | 90.7 | 3.6 |
| July | 93.5 | 3.7 |
| August | 140.5 | 5.5 |
| September | 168.9 | 6.7 |
| October | 196.6 | 7.7 |
| November | 131.6 | 5.2 |
| December | 103.4 | 4.1 |
| Annual Total | 1367.3 | 53.8 |

1. 1Pilson, M.E.Q. 1985. On the residence time of water in Narragansett Bay. *Estuaries* 8:2–14.

   2Narragansett Bay National Estuarine Research Reserve. 2007. An Ecological Profile of the Narragansett Bay National Estuarine Research Reserve. K.B. Raposa and M.L. Schwartz (eds.), *Rhode Island Sea Grant, Narragansett, R.I*. 176pp. [↑](#footnote-ref-1)